## Comment Form
Patient Care Device (PCD) Technical Framework Supplement Public Comment Form
Use the table below to submit comments on 2011 Patient Care Device (PCD) Technical Framework Supplements to the IHE QRPH Technical Committee. Comments submitted by August 1, 2011 will be considered in refining the Trial Implementation versions of these supplements to be published in September 2011. Enter you comments using the table below and send by email to pcdcomments@googlegroups.com. Notes: 1. The 2011 PCD Profile Supplement is: (list below) a. Infusion Pump Event Communication (IPEC) This document is available for download at http://www.ihe.net/Technical_Framework/public_comment.cfm. 2. Designate Volume, Section and Line Number for each comment (or designate as "general"). 3. Issue: Be brief, but do include enough information to indicate the issue of concern to you. 4. Designate Priority of each comment as High (H), Medium (M), Low (L) L: Typo or other minor correction that an editor can manage. Requires no group discussion. M: Medium issue or clarification. Requires discussion, but should not lead to long debate. H: Important issue where there is major issue to be resolved. Requires discussion and debate. 5. Proposal: Any new text or explanation you have. Suggestions for rewording or improvement are greatly encouraged. Having a new proposal makes it easier for committee members to address the issue at hand.
Submitter Name | Supplement Name (click cell and select from list) | Other Document Name | Vol | Section # | Line # | Issue | Proposed Change | Priority (click cell and select from list)
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 105 | A unique event identifier should be mandatory data to (1) link related start, stop, comp together and (2) application acknowledgements for more complex events. The specific fields that comprise the identifier n-tuple should be specified in this TF; hopefully they are the same fields that link ACM PCD-04 alarm transitions. |  | High
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 220 | Lines are missing from Figure X.1-1. |  | Low
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 340 | As a general comment, we should define the content of our event content models so that they can be used by a large number of clinical applications. For example, the 'ethic' of our MDC_PUMP_DELIV_* events should be to convey sufficient information to allow the DOC to determine the drug and how much was delivered at any point in time. In fact, all events should have a description or 'statement of purpose' to make the overall objective of the event very clear. |  | Low
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 398 | Referring to the _START, _STOP, _COMP and other 'tail-end' discriminators listed with the MDC Codes in Table X.1.2-1 -- these should be standardized, and if possible, implemented as standard discriminators in the MDC_ REFIDs. These should correpond to largely standardized 'transitions' or 'phases' of the events. Comment: the IEEE P11073-10102 "Annotated ECG" uses _TPOINT for 'time-point' events and _START and _END for 'interval' events. ACM uses a similar 'start' and 'end' as an enumerated expression of the same concept. |  | Medium
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 416 | I still find it a bit awkward to see each clinical scenario being with _COMP or _STOP and then end with _START; can't you put these in time order, e.g. _START --> (_COMP ¦ _STOP), to match the way it is shown in Figure X.1-1? |  | Medium
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 405 | Tables X.1.2.1-2 and -3 have hyphens in the MDC REFID column headings; they should not, as hyphens are not used for these identifiers. |  | Low
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 416 | In Table X.1.2.1-4, the KVO rate and volume should reflect 'reality', e.g. the rate and amount of fluid being delivered in a consistent manner. If pumps report this differently, then perhaps a flag is required to indicate the KVO reporting policy that is in effect. For example, lets say a SOURCE channel is still delivering a cardiac drug into the line, but the DELIVERY rate is at KVO, resulting in a potentially high concentration of drug in the line that could be delivered to the patient later on when the DELIVERY channel is restarted. |  | High
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 433 | It should be clearly stated in this document (and others) what the event identifier is - is it solely conveyed as an enumerated MDC REFID in OBX-5 of the OBX-3 = MDCX_ATTR_EVT_COND? Or is the MDS, VMD and CHAN that follows required as well? Although the event identifiers that you have proposed in this document are largely unambiguous, many of the event terms in IEEE 11073-10101 are not, and require the MDS, VMD and CHAN to provide the additional context. |  | High
Paul Schluter | Infusion Pump Event Communication (IPEC) |  |  |  | 434 | In your examples, replace '14591346' with a smaller number for clarity (or is there additional meaning embedded in the larger number?) |  | Low
Paul Schluter | Infusion Pump Event Communication (IPEC)

## Supplement Names
Infusion Pump Event Communication
Other

## PriorityNames
High
Medium
Low